## Sheet1
	Issue tracker – imlook4d

	last commit in this sheet					Fatal	Software crashes
	325					Critical	Imlook4d does wrong
						High	Script or model wrong
	next issue number					Medium	Useful
	22					Low	Not very important error, but estethic error
						Very low	Something that really doesn't matter, but should be fixed
						Future	Something good, but leave to future

#	Severity	Status	Fixed in commit #	Description	Who	Date	Solution
1	Medium	Closed	312	Useful to have Threshold within ROI % to be relative max value in ROI		2013-09-30
2	Critical	Open		Mikael Ljungbergs DICOM krashar, pga att samma tag dyker up slumpvis		2013-09-30	Beror på random match at (0008,0008)VL=34
3	Low	Closed	312	PCA at frame 33 if many frames		2013-09-30
4	Critical	Open		Need to reference author of linortfit.m		2013-10-04
5	Critical	Closed	315	Gated PET slices in wrong order.		2013-10-31	Problem was that trig time was not consistent. Made a two pass sorting, sorting slice location for each frame after all other sorting.
6	Critical	Closed	316	ROI info script gives too small volume		2013-11-05
7	Low	Open		Emily lyckas låsa Matlab när hon drar i kanterna på imlook4d-fönstret.		2013-11-18
8		Closed	324	Mac 2013b "Warning slider control can not have a Value outside of Min/Max" över hela skärmen vid start		2013-12-17	Guide high PC slider value was 10000. Fix
9		Closed	324	Mac – Open file *.* does not allow *		2013-12-17	Change to * in fille open dialog for "All files"
10	Critical	Closed	325	With ROI added. Zoom in and click on gray in imlook4d, gives ROI at hidden part of image.		2013-12-17	Fixed, finding Xlim and Ylim

11	Low	Open		Två roi:ar. Hide on/off ger hopp. Cached-image spökar?		2014-01-20
12	Low	Open		imlook4d(file) ger inte något imlook4d_current_handle		15/04/14	Gör Open funktion till Scripts
13	Low	Open		Vore bra att ha ett Loop kommando för fillista		15/04/14	Vet inte riktigt hur det blir bäst.
14	High	Open		ProfileSlices verkar ge fel värden		2014-05-05	Kolla om ROI i varje slice hanteras
15	Medium	Open		Gå tillbaka från Sagital, Coronal
16	Medium	Open		Save DICOM crashes when no network connection		2014-10-03	Pga java kod när jag gör UID från MAC-adress
17	Fatal	Open		Marc could not start imlook4d: matlab timefun duration		2014-11-27	workaround: disable toolbox timefun in path – This is a new thing in Matlab 2014b. I should try to run imlook4d in different (newer) versions of Matlab
18	High	Open		Ref Logan – should implement k2 value ?		2014-12-03
19	Critical	Closed		Slice position tas från optional tag 1041		2014-12-10	Mediso filer får fel ordning i sorting, och i script. Lösning – lägg till slice position i image struct.
19	Medium	Open		Slice location is not updated when yoking two images		2014-12-18
20	Medium	Open		Put in other matrix crashes		2014-12-18
21	Critical	Closed		SUV fails PAT-0001, Åsa B		2015-01-26	Closed, because it worked!